## List.VN
Bach Mai Rehab Center
LIST OF PROPOSED EQUIPMENT FOR PURCHASE
No. | Name of equipment | Name of equipment (Vietnamese) | Remark | Quantity | Unit
1 | Máy nén bóp áp lực hơi toàn thân | A.Yêu cầu chung - Thiết bị mới 100%, sản xuất từ 2022 trở về sau - Thiết bị đạt tiêu chuẩn chất lượng ISO 13485 hoặc tương đương - Điều kiện hoạt động: + Nhiệt độ hoạt động tối đa tới: ≥ 30 độ C + Độ ẩm hoạt động tối đa tới: ≥ 70% B. Yêu cầu cấu hình Máy chính : 01 cái Bộ phụ kiện bao gồm: Ống nén khí xoa bóp chân : 01 đôi Ống nén khí xoa bóp tay : 01 cái Ống nén khí xoa bóp vùng lưng : 01 cái Dây dẫn khí cho chân : 01 cái Dây dẫn khí cho tay : 01 cái Dây dẫn khí cho lưng : 01 cái C. Chỉ tiêu kỹ thuật ≥10 chương trình chế độ Có thể bỏ nén áp lực tại khoang khí bất kỳ Có thể cài đặt nén áp lực từng khoang Có công tắc dừng khẩn cấp Áp suất tối đa: 200mmHg + 20% Thời gian điều trị: 0- 99 phút, điều chỉnh được Nguồn điện: AC 220- 230V, 50~60Hz Công suất tiêu thụ: ≥ 100 VA |  | 02 | bộ
2 | Máy rung toàn thân | 1 máy chính trọng lượng: 55Kg Bảng điều khiển màu 2 tay vịn và trụ giữ 2 bên Bàn rung toàn thân |  | 02 | chiếc
3 | Máy kích thích điện 1 chiều xuyên sọ | Bao gồm: 1 máy kích thích điện, 1 bộ dây dẫn, 2 điện cực, 2 miếng đệm điện cực, 2 dây cao su cố định điẹn cực, 1 sạc điện Bộ kích thích: 71mm x 94mm x 15mm; Trọng lượng: 66g Bộ chương trình: 71mm x 62mm x 15mm; Trọng lượng: 46g Mô-đun lưu trữ: 71mm x 39mm x 15mm; Trọng lượng: 42g Thiết bị sạc: 71mm x 61mm x 15mm; Trọng lượng: 46g Cường độ dòng điện một chiều đều: từ -2000microAmpe đến +2000microAmpe Độ lệch dòng điện: tối đa 2% Offset phần cứng: +/-10 microAmpe Giới hạn điện thế: tối đa +/-16V Tiêu thụ năng lượng: tối đa 0.25W Hoạt động liên tục 90 phút Hiển thị đồ hoạ, một nút ấn | Máy được chỉ định cho đối tượng người bệnh đột quỵ não: liên qiuan đến vận động chi trên, rối loạn ngôn ngữ, mất ngủ , trầm cảm… kết hợp với các phương pháp phục hồi chức năng hỗ trợ cho người bệnh đột quỵ não . | 01 | chiếc
4 | Máy kích thích điện chức năng (chi trên) | Bao gồm: 1 bộ điện cực đa năng có kết nối bluetooth, 1 màn hình điều chỉnh chương trình tập luyện Nguồn điện: Pin 3,7V, 1400mAh Kích thước: 26mm x 76mm x 52mm Trọng lượng: 91gram Công suất tải tối đa: 5000 ohm Công suất đầu ra tối đa: 1,5W Loại xung: hai pha đối xứng hoặc hai pha bù Cường độ xung: độ phân giải 0 – 60mA, 1mA (đối với tải<5000) Độ rộng xung: 150us – 300us, độ phân giải 10us Tần số xung: độ phân giải 1- 40Hz 1Hz Điện áp đầu ra tối đa: 180V Tần số đầu ra tối đa: 1KHz | Hỗ trợ người bệnh tập luyện các chức năng của cổ tay, bàn tay, mở và nắm các ngón tay | 01 | chiếc
5 | Máy kích thích điện chức năng (chi dưới) | Bao gồm: 1 bộ điện cực đa năng có kết nối bluetooth, 1 màn hình điều chỉnh chương trình tập luyện Nguồn điện: Pin 3,7V, 1400mAh Kích thước: 26mm x 76mm x 52mm Trọng lượng: 91gram Công suất tải tối đa: 5000 ohm Công suất đầu ra tối đa: 1,5W Hai mô-đun Bluetooth 3.0 và bluetooth 4.0. chế độ kép Loại xung: hai pha đối xứng hoặc hai pha bù Cường độ xung: độ phân giải 0 – 60mA, 1mA (đối với tải<5000) Độ rộng xung: 150us – 300us, độ phân giải 10us Tần số xung: độ phân giải 1- 40Hz 1Hz Điện áp đầu ra tối đa: 180V Tần số đầu ra tối đa: 1KHz | Hỗ trợ người bệnh tập luyện dáng đi | 01 | chiếc
6 | Máy điện thị liệu (Máy điều trị và kích thích phát âm) | Máy chính : 01 cái, kèm phụ kiện chuẩn sau: Đai thắt điện cực điều trị 01 cái; Công tắc tay 01 cái; Cáp nguồn cung cấp 01 Chiếc; Cáp nối điện cực 01 cái; Điện cực điều trị 04 cái; Đệm điện cực 04 cái; Tài liệu HDSD bảo dưỡng tiếng Anh+ tiếng Việt 01 bộ Điều khiển vi xử lý Số kênh điều trị: 2 kênh.K12 Kiểu luyện tập phát âm: bằng dòng điện tần số thấp và tần số trung. Các dòng điện kích thích phát âm: G, TENS, UR, IG30, IG50, FM, FaS, T/R, AMF, MT, KOTS Công suất phát: 20mA / 500 ohms ; G = 5mA / 500 ohms Chương trình chẩn đoán: kiểm tra mức độ tổn thương của bệnh nhân Điều khiển máy: - Núm xoay vô cực đa hướng - Công tắc tay Hiển thị các thông số cài đặt khi điều trị: kiểu cảm ứng màu Khả năng lưu trữ: Có, một số phác đồ điều trị do người sử dụng lưu vào Dụng cụ kết hợp điều trị: Có, loa phát tiếng các bài luyện âm, luyện nói Khả năng hiển thị:Hình ảnh 3D minh hoạ vùng sẽ điều trị và vị trí đặt điện cực. Tự động dừng phát xung điện khi dòng ra ở điện cực bị gián đoạn. Thời gian điều trị: từ 0,5 đến 60 phút, nấc chỉnh 0,5 phút. Hỗ trợ thông tin kết quả điều trị trên điện thoại thông minh | Hỗ trợ chẩn đoán và điều trị người bệnh phát âm nói và tập nuốt | 01 | chiếc
7 | Máy Biofeedback tập nuốt cầm tay | Bao gồm: 1 máy chính, 1 bộ cáp điện cực, 1 bộ phụ kiện chuẩn Cấu hình tiêu chuẩn: - Máy chính màn hình màu cảm ứng: 01 máy - Cáp điện cực điều trị khó nuốt: 04 cái - Cáp điện cực tham chiếu: 01 cái - Công tắc: 01 cái - Vỏ cao su chống sốc: 01 cái - Đế kê máy: 01 cái - Điện cực dán: 01 gói - Điện cực tham chiếu: 01 - Pin AAA 1,5V: 04 viên - Hướng dẫn sử dụng Tiếng Việt: 01 quyển |  | 01 | chiếc
8 | Máy kích thích TENS/ EMS cầm tay | Máy chính: 01 cái Vali đựng máy: 01 cái Cáp điện cực: 01 bộ Tấm điện cực dán: 01 bộ Số kênh: 2 kênh, điều chỉnh độc lập giữa hai kênh Xung: có thể điều chỉnh từ 0-110mA đỉnh 500 ohm mỗi kênh Dạng sóng: Xung vuông Bi-Phasic không đối xứng. Hiệu điện thế: 0-55V( tải 500ohm) Nguồn cấp: Pin 9V Kích thước: 11,8x 6x 3,1cm Trọng lượng: 157g bao gồm cả pin Tỉ lệ xung: Điều chỉnh 2-150 Hz, 1Hz mỗi bước Độ rộng xung: Điều chỉnh từ 50- 300 micro giây, chỉnh 10 micro giây mỗi bước Thời gian hoạt động: Điều chỉnh 2-90 giây, 1 giây mỗi bước chỉnh Ramp Time: chỉnh 1-8 giây, 1 giây mỗi bước, tăng và giảm trong cài đặt giá trị Kiểu dòng điện: 6 kiểu dòng TEN: B, N, M, S1, S2 và P 3 kiểu dòng EMS: C, S, A B( Burst mode): Tỉ số Burst điều chỉnh 0.5 đến 5Hz, độ rộng xung 50- 300 µs, tần số cố định 100Hz N( Normal mode): tỉ số và độ rộng xung điều chỉnh được, chạy liên tục dựa trên thông số cài đặt M( modulation mode): Kiểu kết hợp điều biến tỉ lệ và độ rộng xung. S1( S1 mode): Độ rộng xung điều chỉnh tự động trên danh nghĩa 10s mỗi chu kỳ. S2( S2 mode): Tương tự S1 giá trị cài đặt cao hơn 70% so với S1 là 40% C( Constant mode): Điều chỉnh tần số và độ rộng xung. |  | 04 | chiếc

## List.EN
Bach Mai Rehab Center
LIST OF PROPOSED EQUIPMENT FOR PURCHASE
No. | Name of equipment | Name of equipment (English) | Remark | Quantity | Unit
1 | Air compression machine for full body | General requirement: - 100% new, manufactured from 2022 onwards - ISO 13485 quality standards or equivalent - Operating conditions: + Maximum operating temperature: ≥ 30 degrees Celsius + Maximum operating humidity: ≥ 70% Supply requirement: - Main device: 01 pce - Accessory set includes: + Leg cuff: 01 pair + Arm cuff: 01 pce + Waist cuff: 01 pce + Hose for leg cuff: 01 pce + Hose for arm cuff: 01 pce + Hose for waist cuff: 01 pce Technical specifications: - ≥10 modes program - Can skip pressure at any chamber; Pressure value of each chamber can be set independently; There is an emergency stop switch - Maximum pressure: 200mmHg + 20% - Adjustable time: 0- 99 minutes - Rate voltage: AC 220- 230V, 50~60Hz - Power consumption: ≥ 100 VA |  | 02 | set
2 | Whole body vibrator machine | Include 1 main machine 1 main machine weight: 55Kg Color control panel 2 handrails and pillars on both sides Full body vibrating table |  | 02 | piece
3 | Transcranial direct current stimulation (tDCS) | Include: 1 electrical stimulator, 1 set of wires, 2 electrodes, 2 electrode pads, 2 electrode fixing rubber bands, 1 electric charger: - Stimulator: 71mm x 94mm x 15mm; Weight: 66g - Program set: 71mm x 62mm x 15mm; Weight: 46g - Storage module: 71mm x 39mm x 15mm; Weight: 42g - Charging device: 71mm x 61mm x 15mm; Weight: 46g - One way direct electric current intensity: all from -2000 microAmpe to +2000 microAmpe - Electric current deviation/difference: 2% maximum - Hardware offset: +/-10 micro Ampe - Voltage limit: maximum +/-16V - Power consumption: 0.25W maximum - Operates continuously for 90 minutes - Graphical display, one pressing button | - The machine is indicated for stroke patients: related to upper limb movement, language disorders, insomnia, depression... combined with rehabilitation methods to support stroke patients. | 01 | piece
4 | Functional electrical stimulation (upper limbs) | Include: 1 multi-function electrode set with bluetooth connection, 1 screen to adjust the exercise/pratising program Power supply: 3.7V, 1400mAh battery Dimensions: 26mm x 76mm x 52mm Weight: 91grams Maximum load capacity: 5000 ohm Maximum output power: 1.5W Pulse type: symmetrical two-phase or 2 compensate capacitors Pulse intensity: resolution 0 – 60mA, 1mA (for load<5000) Pulse width: 150us – 300us, 10us resolution Pulse frequency: resolution 1- 40Hz 1Hz Maximum output voltage: 180V Maximum output frequency: 1KHz | Support the patient in practicing the functions of the wrist, hand, opening and closing the fingers | 01 | piece
5 | Functional electrical stimulation ((lower limbs) | Include: 1 multi-function electrode set with bluetooth connection, 1 screen to adjust the exercise/pratising program Power supply: 3.7V, 1400mAh battery Dimensions: 26mm x 76mm x 52mm Weight: 91grams Maximum load capacity: 5000 ohm Maximum output power: 1.5W Two modules Bluetooth 3.0 and bluetooth 4.0. dual mode Pulse type: symmetrical two-phase or 2 compensate capacitors Pulse intensity: resolution 0 – 60mA, 1mA (for load<5000) Pulse width: 150us – 300us, 10us resolution Pulse frequency: resolution 1- 40Hz 1Hz Maximum output voltage: 180V Maximum output frequency: 1KHz | Support patients with pratising gait /walking | 01 | piece
6 | Electrotherapy unit (treatment of patients with swallowing and voice disorders by stimulation currents) | 1 Main unit incl. at least standard accessories as below: Elastic fabric tape 1 pc; Hand button 1 pc; Power cable 1 pc; Patient cable 1 pc; Plate electrode 4 pcs; Viscose bag 4 pcs; Instructions for use (English-Vietnamese) 1 pc. -Microprocessor control -Number of treatment channels: 2 K12 channels. -Pronunciation training type: by low frequency and mid frequency. -Vocal stimulation electric currents: G, TENS, UR, IG30, IG50, FM, FaS, T/R, AMF, MT, KOTS -Output power: 20mA / 500 ohms ; G = 5mA / 500 ohms -Diagnostic program: check the patient's injury level -Machine control: - Multi-directional infinity knob; - Hand switch -Display of treatment settings: color touch type -Storage capacity: Yes, some treatment regimens can be saved by the user -Treatment combination tools: Yes, speakers play pronunciation and speaking exercises -Visibility: 3D images illustrate the area to be treated and electrode placement. -Automatically stops generating electric pulses when the output current at the electrode is interrupted. -Treatment time: from 0.5 to 60 minutes, adjustable level: 0.5 minutes. Support treatment results information on smartphones | The purpose of device is treatment of patients with swallowing and voice disorders by stimulation currents. | 01 | piece
7 | Biofeedback and neuromuscular electrical stimulation | Includes: 1 main machine, 1 set of electrode cables, 1 set of standard accessories Standard configuration: - Main machine with color touch screen: 01 machine - Electrode cables for dysphagia treatment: 04 pieces - Reference electrode cable: 01 piece - Switch: 01 piece - Shock-proof rubber case: 01 piece - Machine stand: 01 piece - Paste electrode: 01 package - Reference electrode: 01 - 1.5V AAA batteries: 04 pieces - User manual (Vietnamese): 01 book |  | 01 | piece
8 | TENS/EMS combo stimulation | Main machine: 01 pc Suitcase for carrying machine: 01 piece Electrode cable: 01 set Sticky electrode plate: 01 set Channel: dual, isolated between channels Pulse Amplitude: adjustable, 0-110mA peak into 500 ohm load each channel. Wave form: Asymmetrical Bi-Phasic square pulse. Voltage: 0-55V( load: 500ohm) Power source: One 9 Volt Battery Size: 11,8(L)x 6cm (W)x 3,1cm (H) Weight: 157grams with battery Pulse rate: adjustable from 2 to 150 Hz, 1Hz/ step Pulse Width: Adjustable from 50- 300 microseconds, 10 µs /step On Time: Adjustable 2-90 seconds, 1 Sec./step Off Time: Adjustable 2-90 seconds, 1 Sec./step Ramp Time: Adjustable 1-8 seconds, 1 Sec./ step, The "On" time will increase and decrease in the settting value. Mode: Six TENS modes: B(burst), N (normal), M (modulation rate & width), S1 (Modulation Width), S2 (Modulation Width) and P three EMS modes: C (constant), S (synchronous), A (Alternate) B( Burst mode): Burst rate: adjustable 0.5 ~ 5Hz, pulse width adjustable 50- 300 µs, Frequency fixed = 100Hz N( Normal mode): the pulse rate and pulse width are adjustable, it generates continous stimulation based on the setting value. M( modulation mode): modulation mode is a combination of pulse rate and pulse width modulation. S1 (S1 mode): pulse width is automaticaaly varied in a cyclic pattern over a nominal 10 second period. S2 (S2 mode): same as S1 but pulse width decrease over a period of 4 seconds from the initial setting to a value 70% less. C (Constant mode): The pulse rate and pulse width are adjustable. |  | 04 | piece